## Sheet1
SARS-COV-2 Dual Ab assay (IgM, IgG) rapid diagnostic test kits : Option appraisal from data available 21/3/20
Manufacturer | Kit | Overview/Technology | CE Mark | Eye/Electronic | Intended use | Sample type | Sample storage | Sample volume | Effect of haemolysis | Pre-analytics | Assay time | Max time to read | Reagent storage | Shelf life | Diagnostic sensitivity | Diagnostic specificity | Interferences | Cross-reactivity | Imprecision | Limitations
AmonMed Biotechnology Co. Ltd | COVID-19 IgM/IgG test kit (Rare Earth Nano Fluorescence Immunochromatography) (CE-IVD) | Four different Dry Fluorescent immunoassay analysers available for a range of workloads. No mention of the COVID-19 assay on the website. | Yes
| COVID-19 IgM/IgG test kit (Colloidal Gold) (CE-IVD) | No mention of this COVID-19 assay on the website. | Yes
Beijing Abace Biology Co. Ltd. | COVID-19 Antibody (IgG/IgM)Test Kit (Colloidal Gold Immunochromatography) (CE-IVD) | Unable to trace company | Yes
Beijing Kewei Clinical Diagnostic Reagent Inc. | Kewei COVID-19 IgG/IgM Fluorescence Rapid Test Kit (CE-IVD) | No mention of COVID-19 assay on website | Yes
BioMedomics, Inc. Distributed in the UK by AphaBioLabs (currently out of stock & inviting pre-orders) | COVID-19 IgM-IgG Dual Antibody Rapid Test (CE-IVD) | Colloidal gold immunochromatography, IgM, IgG and C lines. Emailed AlphBioLabs for info. | Yes | Eye | Healthcare Professionals, qualitative testing | Fresh Serum, Plasma, WB (capillary or venous) EDTA, Heparin, Na Citrate anticoags. Small study with 7 pos and 3 neg samples showed 100% concordance for test results with venous blood. | Serum/Plasma 2-8°C for 5 days Anticoagulated WB 2-8°C for 7 days | 20 ul WB, 10 ul serum/plasma | Not mentioned | Apply specimen from capillary sampler into test sample well. Add 2-3 drops of buffer solution from bottle. | 10 minutes | 15 minutes | In dark 18-26°C | 12 months | The patients recruited were tested by RT-PCR and assessed clinically. These samples were collected from various hospitals and CDC testing laboratories (total 8) at 6 different provinces in China. Venous blood, 525 cases (397 Pos, 128 Neg) Sensitivity 88.66%. A sub-set of 58 patients tested day 8 - day 33 from first symptoms gave sensitivity 100%. Data on time from first symptoms was not collected for the rest of the groups. | The patients recruited were tested by RT-PCR and assessed clinically. These samples were collected from various hospitals and CDC testing laboratories (total 8) at 6 different provinces in China. Venous blood, 525 cases (397 Pos, 128 Neg) Specificity 90.63%. Specificity data not provided for sub-set of 58 patients tested day 8 - day 33 from first symptoms. | Data not presented | Does not cross react with Flu A, Flu B or RSV Abs. SARS Cov1 has not been tested. | Data not presented | This test has not been reviewed by FDA, Negative results do not rule out SARS-CoV-2 infection, particularly in those who have been in contact with the virus. Follow-up testing with a molecular diagnostic should be considered to rule out infection in these individuals. Results from antibody testing should not be used as the sole basis to diagnose or exclude SARS-CoV-2 infection or to inform infection status. Positive results may be due to past or present infection with non-SARS-CoV-2 coronavirus strains, such as coronavirus HKU1, NL63, OC43, or 229E.
Core Technology Co. Ltd | COVID-19 IgM/IgG Ab Test (CE-IVD) | No mention of COVID-19 assay on website | Yes
CTK Biotech Inc. | OnSite COVID-19 IgG/IgM Rapid Test (CE-IVD) | Limited details on website. Emailed for technical details. | Yes | Eye. But also compatible with CTK rapid test reader. |  | WB/serum/plasma |  |  |  |  | 10 minutes |  |  |  | 96.9% | 99.4%
Diagreat | Lateral flow fluorescence immunoassay | Separate IgG and IgM tests but fluorescence so machine read only. Emailed for further details. | Yes | Single channel POC reader |  | WB/serum/plasma |  |  |  | Apply specimen from capillary sampler into test sample well. Add drops of buffer solution from bottle. | 15 minutes |  | 2-30°C | 18 months
Dynamiker Biotechnology (Tianjin) Co. Ltd. | 2019 nCOV IgG/IgM Rapid Test (CE-IVD) | Emailed for more info - may only be available in Afghanistan, China and America | Yes | [12 empty] | one subset data of 58 patients in Wuhan, it was found that 94.83% of the positive patients had both IgM and IgG positive test lines, and 1.72%, 3.45% had only IgM or only IgG positive lines respectively (table 2). The test time was at day 8 to day 33 after infection symptoms appeared
GenBody Inc. | GenBody COVID-19 IgM/IgG (CE-IVD) | Website says COVID19 coming soon | Yes
Getein Biotech, Inc., | One Step Test for Novel Coronavirus (2019-nCoV) IgM/IgG antibody (Colloidal Gold) (CE-IVD) | Product advertised on website but no additional information | Yes
Hangzhou Biotest Biotech Co,Ltd. | COVID-19 IgG/IgM Rapid Test Cassette (Whole Blood/Serum/Plasma) (CE-IVD) | Rightsign Biotest assay. Limited info on website but tables with performance data are identical to those given with the Nal Von Minden test kit. | Yes
Innovita Biological Technology Co. Ltd, | 2019-nCoV Ab Test (Colloidal Gold) (IgM/IgG Whole Blood/Serum/Plasma Combo) (CE-IVD) | Limited information on website | Yes
InTec Products, Inc. | Rapid SARS-CoV-2 Antibody (IgM/IgG) Test (CE-IVD) | Nothing obvious on line | Yes
Liming Bio-Products Co., Ltd, | COVID-19 IgG/IgM Combo Rapid Test Device (CE-IVD) | Limited info on website | Yes | Eye |  | WB/serum/plasma |  |  |  |  | 15 minutes |  |  |  | IgG sensitivity 82.7% IgM sensitivity 64.7% Data from 173 positive subjects and 33 negative subjects | IgG specificity 100% IgM specificity 100% Data from 173 positive subjects and 33 negative subjects
MedicalSystem Biotechnology Co., Ltd., | COVID-19 IgM/IgG Rapid Test Cassette (CE-IVD) | No mention of COVID-19 assay on website | Yes
SD Biosenor Inc | STANDARD Q COVID-19 IgM/IgG Duo Test (CE-IVD) | Two separate IgG and IgM tests | Yes | Eye
SensingSelf, Pte, Ltd, Singapore, | EDR COVID 19 Rapid Test Kit (IgM/IgG) (CE-IVD) | Limited information on website | Yes | Eye
Servoprax GmbH | Cleartest Corona, Covid-19 (CE-IVD) | No mention of COVID-19 assay on website | Yes
Shenzhen Bioeasy Biotechnology Co., Ltd., | Novel Coronavirus (2019-nCoV) IgG/IgM detection kit (colloidal gold immunochromatography) (CE-IVD) | Videos of Bioeasy test on website but no technical info | Yes
Sugentech, Inc. | SGTi-flex COVID-19 IgM/IgG (CE-IVD) | Limited information on website | Yes
Sure BioTech USA Co Ltd | SARS-CoV-2 IgM/IgG Ab Rapid Test (CE-IVD) | Separate IgM and igG test. No technical data on website. | Yes | Eye |  |  |  | 10ul |  |  | 15 minutes
Tianjin MNCHIP Technologies Co., Ltd., | Anti-COVID-19 virus IgM/IgG rapid test kit (Colloidal gold assay) (CE-IVD) | No mention of COVID19 test on website but they have POC readers | Yes | Eye/reader
VivaChek Biotech (Hangzhou) Co., Ltd | VivaDiag COVID-19 IgM/IgG Rapid Test (CE-IVD) Contact | Vivadiag colloidal gold. No evaluation data on website. | Yes | Eye | Healthcare professionals only | WB/serum/plasma | 2-30°C | 10ul |  | Apply specimen from capillary sampler into test sample well. Add 2 drops of buffer solution from bottle. 18-25°C | 15 minutes | 20 minutes |  | 12 months
Xiamen Biotime Biotechnology Co., Ltd., | SARS-CoV-2 IgG/IgM Rapid Qualitative Test Kit (CE-IVD) | No technical data on website | Yes
Una Health | Nal Von Minden NADAL(R) COVID-19 IgG/IgM Test | Lateral flow immunochromatography, IgM, IgG and C lines | Yes | Eye | Professional use only, qualitative testing | Fresh Serum, Plasma, WB (capillary or venous)K EDTA,Li/Na Heparin, Na/K Citrate, Na Oxalate anticoags. Anti-coagulant equivalence studies were only performed on SARS Cov 2 negative samples (100% concordance). | Do not leave specimens at room temp for long periods of time. Serum/Plasma 2-8°C for 3 days Anticoagulated WB 2-8°C for 2 days | 10 ul WB/serum/plasma | Icteric, lipaemic or haemolysed specimens may give inaccurate result. Interference studies indicate no effect from Hb at 10 g/l. | Bring tests, specimens and buffer to room temp. Apply specimen from capillary sampler into test sample well. Add 2 drops of buffer solution from bottle. | 10 minutes | 20 minutes | 2-30°C | Not stated | Assay performance was assessed by comparing test kit results to clinical diagnoses confirmed by PCR. The group was made up of 85 positive cases and 371 suspected/healthy cases. NB: "Early state" samples (those taken "shortly" after the onset of first symptoms) were excluded. Discrepancies were observed here vs PCR as expected. IgG sensitivity 99.9% IgM sensitivity 91.8% Combined sensitivity 94.1% IgG Accuracy 99.6% IgM Accuracy 97.8% Combined Accuracy 98.2% | Assay performance was assessed by comparing test kit results to clinical diagnoses confirmed by PCR. The group was made up of 85 positive cases and 371 suspected/healthy cases. NB: "Early state" samples (those taken "shortly" after the onset of first symptoms) were excluded. Discrepancies were observed here vs PCR as expected. IgG specificity 99.5% IgM specificity 99.2% Combined specificity 99.2% IgM Accuracy 99.6% IgG Accuracy 97.8% Combined Accuracy 98.2% | Interference studies performed for Acetaminophen, Acetyl salicylic acid, Albumin, Ascorbic acid, Bilirubin, Caffeine, Creatine, Ethanol, Gentisic acid, Hb, Oxalic acid, Uric acid. No interference observed at the concentrations tested. | Cross reacts with anti-SARS Cov and some samples with Rheumatoid Factor. MERS Cov cross reactivity is not ruled out. Does not cross react with Abs to Flu A, Flu B, RSV, Adenovirus, HBSAg, Syphillis, H. Pylori, HIV, HAMA. | Repeatability studies with 10 replicates across 3 test lot numbers for positive and negative samples showed 100% agreement. | At the beginning of the disease, the concentration of anti-SARS-CoV-2 IgM may be below the detection limit of the SARS-CoV-2 IgM may be below the detection limit of the test. The continued presence or absence of antibodies cannot be used to determine the success or failure of therapy. Results from immunosuppressed patients should be interpreted with caution. A positive test result can also occur in case of negative PCR results because antibodies are still present in the blood after the illness and can be detected. A negative result does not at any time preclude the possibility of a SARS-CoV-2 infection.
Biospace | 20/20 Bioresponse CoronaCheck ™ | Website suggests sales may be USA only | Yes | [11 empty] | CoronaCheck was evaluated in China using 760 clinical samples. In this cohort, 286 samples were confirmed positive. Manufacturers report a sensitivity of 97.2% and specificity of 92%. | CoronaCheck was evaluated in China using 760 clinical samples. In this cohort, 286 samples were confirmed positive. Manufacturers report a sensitivity of 97.2% and specificity of 92%.
BioPanda Reagents Ltd, Belfast | COVID-19 IgM/IgG | Shipping to UK only. Emailed for details. | No mention | Eye |  | WB/serum/plasma
SureScreen Diagnostics, Derby, UK | COVID-19 IgM/IgG | Limited data on website. Emailed for details. Public Health England said to avoid using this test because there is 'little information on the accuracy of the test'. | Yes | Eye | Professional use only | WB/serum/plasma |  |  |  |  | 10 minutes |  |  |  | 181 samples tested vs PCR- no detail IgG sensitivity 100% Accuracy 99.6% IgM sensitivity 91.8% Accuracy 97.8% | 181 samples tested vs PCR - no detail IgG specificity 91.8% Accuracy 99.6% IgM specificity 99.2% Accuracy 97.8%
Bioconnections, Stoke on Trent | Separate IgG and IgM kit from Era Biology Genobio Pharmaceuticals Ltd) | Lateral flow "dipstick " test. Requires microtitre plate and predilution of samples. Company issued evaluation paper that did not pertain to this product. Not considered further. | Yes | Eye
Healgen (Distributed through Menarini Diagnostics) | COVID-19 IgG/IgM | Lateral Flow Colloidal Gold IgG, IgM, Control Line | Yes | Eye | Professional use | WB/serum/plasma |  | 5 uL |  | N/A | 2-10 minutes | 10 minutes | 2-30 Celsius | 22 months | IgM 87.9%, IgG 97.2 % | 100% for IgG and IgM
EcoTest (Matrix Diagnostics | COVID-19 IgG/IgM | Lateral Flow | Yes | eye | Professional use only | WB/serum/plasma | serum/plasms 2-8C for 7 days, whole blood refrigerated up to 3 days. Long term storage of serum and plasma at -20C | 5ul |  |  | read at 15 minutes | 15 minutes | 2-30C |  | IgM 93.7%, IgG 98.8% | 99.1% IgM, 98.7% IgG
Distributer:LumiraDx (LumiraTek) Kit made by Nagzhou Biotest Biotech Ltd. | COVID-19 IgG/IgM | Lateral Flow | Yes | eye |  | WB/serum/plasma |  | 10ul |  |  | Read at 10 minutes |  | 2-30C |  | IgM 91.8%, IgG 100% | IgM 99.5%, IgG 99.2%
EDAN - | Colloidal Gold-labelled. SARS-CoV-2 IgM&IgG Rapid Test | Lateral Flow | Yes | eye |  | WB/serum/plasma |  | 10-20ul |  |  | 5 - 10 minutes | 10 minutes | 2-30C | 18 months